## Command List
Name | Notes | Direction | Parameters (Separated by commas in GECP protocol)
Identify | Requests the Device Name and Firmware version NOTE: This command can be sent before the instrument has been Locked | COMMAND OUTPUT | Device ID Format: "%s" Example: "VERITY 3011 PUMP" | Firmware Version Format: %u.%u.%u.%u Example: "1.0.0.0"
Get Device ID | Requests the Device Name and Firmware version NOTE: This command can be sent before the instrument has been Locked NOTE: Sending to destination ID "1" returns the controller board's device ID. Sending to destination ID "2" returns the pump board's device ID. | COMMAND OUTPUT | Device ID Format: "%s" Example: : "VERITY 3011 PUMP CONTROLLER" "VERITY PUMP MODULE" | Firmware Version Format: %u.%u.%u.%u Example: "1.1.0.13"
Lock | Lock the touch display so that the instrument is only operable by a remote host over the rear comms. | COMMAND INPUT
Lock With Options | Lock the touch display so that the instrument is only operable by a remote host over the serial comms. Title is displayed on the lock screen. Pump Status Messages and Rate determine if status messages are set out the serial port and the time between messages. | COMMAND INPUT | Title Format: "%s" Example: "VERITY 3011" | Pump Status Messages Format: "%s" Range: "On", "Off" | Pump Status Message Rate Units: milliseconds Format: "%u" Range: 1000 - 60000
Unlock | Unlock the touch display so that the instrument is only operable by the touch display. | COMMAND INPUT
Clear Error | Instructs the instrument to clear errors specified in the parameters list. | COMMAND INPUT | Mode Format: "%s" String Values: "ALL" clears current Error and buffered errors "LOG" clears the command log "" clears current Error only
Get Error | Requests a report of the existing errors. Read the current error reported; No error Command,Error State,Error Number/Total Number of Errors,Event/Command,Return Code,Error String Example: ?[4,2,0,RSP,0,3(Get Error,No Error,,,,Error log empty)]? ?[4,2,0,RSP,0,3(Get Error,In Error,1/1,Event,18,Maximum Pressure Limit)]? ?[5,2,0,RSP,0,3(Get Error,No Error,1/1,Set Pump Refill Time,15,Invalid Refill Time)]? | COMMAND OUTPUT | Instrument State Format: "%s" Example: "State¦No error" | Error Information Format: "Error 1/%u¦Cmd=%s" Example: "Error 1/1¦Cmd=Set Pump Flow Rate¦RtnCode = 11¦Err = Invalid Flowrate" "Error log empty"
Set Error | Command to add an error to the error log and start the status LED blinking. | COMMAND INPUT | Error Event Return Code Format: "%u" Range: 6 - 255 | Error Message Format: "%s" Example: "Pressure Sensor Lost"
Send Pump Run Status | Gets a Pump Run Status message. | COMMAND OUTPUT | Current Flow Rate Units: mL/min Format: "%.3f" Example: "10.000" | Current Pressure Units: bar Format: "%.2f" Example: "500.25" | Current Solvent Level Units: mL Format: "%.3f" Example: "10000.25" | Current Run State Format: "%s" String Values: "Running" "Run Complete" "Paused" "Stopped" | Current Error Condition Format: "%s" String Values: "High Pressure" "High Pressure Method" "Over 75 Pressure" "Low Pressure" "Low Pressure Method" "Low Volume" "Input 4 Method" "Normal" | Method Name Format: "%s" Range: 16 characters | Current Method Loop Format: "%u" Range: 1 - 10000 | Current Method Event Number Format: "%u" Range: 1 - 50 | Pause Timer Time Units: minutes Format: "%.3f" Example: "5.314" | Current Timestamp Units: minutes Format: "%.3f" Example: "5.314" | Ramp Time Units: minutes Format: "%.3f" Example: "5.314" | Method Loop Time Units: minutes Format: "%.3f" Example: "5.314" | Current Dispense Time or Volume Units: minutes or mL Format: "%.3f" Example: "5.314" | Operation Hours Units: hours Format "%.2f" Example: "5.34"
Set Alarm | Instructs the instrument to to the buzzer alarm off or on. | COMMAND INPUT | Buzzer Format: "%s" Range: "On", "Off"
Get Pump Flow Rate | Requests the current flow rate of the pump. | COMMAND OUTPUT | Current Flow Rate Units: mL/min Format: "%.3f" Example string "1.234"
Set Pump Flow Rate | Instructs the pump to start at the specified flow rate (in mL/min) The Flow rate is bounded based on the limits of the configured pump head. Subsequent Set Pump Flow Rate will modify any existing running Set Pump Flow Rate command with the new parameters. If in another mode (Dispense by Volume, Dispense by Time), the command will be rejected. ?[2,1,0,RSP,0,9(Set Pump Flow Rate,Invalid Mode Change Requested)]? | COMMAND INPUT | Flow Rate Units: mL/min Format: "%.3f" Range: 5 SS, 0.001 - 5.000 10 SS, 0.050 - 10.000 | Acceleration Units: RPM/s Format: "%u" Range: 1 - 700 OPTIONAL | Current Units: percentage Format: "%u" Range: 1 - 100 Notes: Maximum value can be constrained through the used of NV-RAM settings. OPTIONAL | PFD Units: 8-bit DAC values Format: "%u" Range: 0 - 255 Notes: 255- 192 Fast Decay Mode 191-128 Mid-Fast Decay Mode 127-0 Slow Decay Mode Mid-Slow Decay mode not used OPTIONAL
Dispense by Volume | Instructs the pump to dispense a unit of volume. Subsequent Dispense by Volume will modify the current flow rate parameter but continue with the current dispensed volume. Typing a dispense volume smaller than what has already been dispensed will immediately stop the pump. If in another mode (Flow, Dispense by Time), the command will be rejected. | COMMAND INPUT | Flow Rate Units: mL/min Format: "%.3f" Range: 5 SS, 0.001 - 5.000 10 SS, 0.050 - 10.000 | Volume to Dispense Units: mL Format: "%.3f" Range: unchecked
Dispense by Time | Instructs the pump to dispense for a period of time. Subsequent Dispense by Time will modify any existing running Dispense by Time command with the new parameters. If in another mode (Flow, Dispense by Volume) the command will be rejected. Set Flow Rate or Set Dispense by Time with a flow rate of 0mL/min will stop the run. | COMMAND INPUT | Flow Rate Units: mL/min Format: "%.3f" Range: 5 SS, 0.001 - 5.000 10 SS, 0.050 - 10.000 | Duration Units: minutes Format: "%.2f" Range: unchecked
Get Dispense Volume | Gets the total volume dispensed during a dispense. Reports the dispensed volume (in mL) for the current Dispense by Volume run. When the current run is stopped, the dispensed volume will return zero. | COMMAND OUTPUT | Dispensed Volume Units: mL Format: "%.3f" Example: "5.314" | Total Volume to Dispense Units: mL Format: "%.2f" Example: "5.35"
Get Pump Head | Requests the current pump head configuration setting. | COMMAND OUTPUT | Pump Head Format: "%s" Example: "5 SS" | Minimum Flow Rate Units: mL/min Format: "%.3f" Example: 0.001 | Maximum Flow Rate Units: mL/min Format: "%.3f" Example: 5.000 | Minimum Pressure Units: bar Format: "%u" Example: 1 | Maximum Pressure Units: bar Format: "%u" Example: 600
Set Pump Head | Command to set the pump configuration to the selected pump head. Value is entered into the NV-RAM. Reports the status of the "Set Pump Head" command. | COMMAND INPUT | Pump Head Format: "%s" Example: "5 SS"
Get Pump Refill Time | Requests the Refill Time from the pump. | COMMAND OUTPUT | Refill Time Units: seconds Format: "%.3f" Example response: "0.125"
Set Pump Refill Time | Sets the pump refill time in seconds. Range is 0.125 - 1.00 seconds. | COMMAND INPUT | Refill Time Units: seconds Format: "%.3f" Example response: "0.125"
Get Compressibility | Example response "Methanol,121, 119, 0.105938,874.101344" Parameters returned in order: "solution name, compressibility factor @ zero bar, compressibility factor @ actual pressure, coefficient One, coefficient Two". | COMMAND OUTPUT | Solution Name Format: "%s" | Compressibility Factor Zero Units: mBar-1 Format: "%u" Notes: This is the compressibility at 0 bar | Compressibility Factor Units: mBar-1 Format: "%u" Notes: This is the compressibility at currently measured pressure | Compressiblity Coefficient One Format: "%.6f" | Compressiblity Coefficient Two Format: "%.6f"
Set Compressibility | Reports the status of the "Set Compressibility" command | COMMAND INPUT | Liquid Index Format: "%u" Range: 0 - 14 Legend: 0 Other Liquid 1 Water 2 Methanol 3 Acetonitrile 4 Tetrahydrofuran 5 Benzene 6 n-Hexane 7 Carbon Tetrachloride 8 Water-Methanol, 50-50 9 Chloroform 10 Methylene Chloride 11 Ethanol 12 Acetone 13 n-Heptane 14 Diethyl Ether | Compressibility Term Beta Units: 1E-6 bar Format: "%u" Range: 0 - 2000 Notes: only valid with liqud index of zero OPTIONAL
Set Compressibility Coeff | Reports the status of the "Set Compressibility Coeff" command | COMMAND INPUT | Liquid Index Format: "%u" Range: 0 - 14 Legend: 0 Other Liquid 1 Water 2 Methanol 3 Acetonitrile 4 Tetrahydrofuran 5 Benzene 6 n-Hexane 7 Carbon Tetrachloride 8 Water-Methanol, 50-50 9 Chloroform 10 Methylene Chloride 11 Ethanol 12 Acetone 13 n-Heptane 14 Diethyl Ether | Compressibility Coefficient A Format: "%.6f" Range: 0 - 10000 Notes: only valid with a liquid index of zero OPTIONAL | Compressibility Coefficient B Format: "%.6f" Range: 0 - 10000 Notes: If Compressibilty Coefficient A is equal to zero then this coefficient is forced to be zero as well If the term is not supplied then the value for water is used for this term. Water value for coefficient B is 3771.006 OPTIONAL
Get Output Contacts | Requests the states of the instrument's outputs | COMMAND OUTPUT | +24V Output Format: "%s" Range: "Open", "Closed" | Relay Format: "%s" Range: "Open", "Closed"
Output Contacts | Instructs the instrument to control the output contacts. | COMMAND INPUT | +24V Output Format: "%s" Range: "Open" "Closed" "" | Relay 1 Format: "%s" Range: "Open" "Closed" ""
Get Input Contacts | Requests the states of the digital inputs. | COMMAND OUTPUT | Start/Stop Format: "%s" Range: "Open", "Closed" | Pause/Resume Format: "%s" Range: "Open", "Closed" | Program Wait Format: "%s" Range: "Open", "Closed" | Error Format: "%s" Range: "Open", "Closed" | Power Down Format: "%s" Range: "Open", "Closed"
Get Pressure | Request the current system pressure. | COMMAND OUTPUT | Pressure Units: bar Format: "%.2f" Example string "123.45"
Start Pressure Samples | Instructs the instrument to start transmitting pressure messages at the defined rates. | COMMAND INPUT | Sample Interval Units: ms Format: "%u" Range: 200 - 65000 | Samples Per Message Format: "%u" Range: 1 - 10
Stop Pressure Samples | Instructs the instrument to stop transmitting pressure messages. | COMMAND INPUT
Set Inlet Pressure | Set the inlet pressure variable to be used in the compressibility computation. | COMMAND INPUT | Inlet Pressure Units: bar Format: "%.2f" Range: 0 - 100
Get Inlet Pressure | Returns the value of the inlet pressure setting. | COMMAND OUTPUT | Inlet Pressure Units: bar Format: "%.2f" Range: 0 - 100
Home | Home command is used on power up and can be user sent to index the motor to a known position. The successful home position is based on the mechanical index and then a precharge movement is completed to precompress the solution. Reports the status of the home command at the completion of the movement. | COMMAND OUTPUT | Home State Format: "%s" Example: "Homed"
Get Motor Status | Request the status of the pump motor. | COMMAND OUTPUT | Motor Status Format: "%s" Legend: Error: high pressure error or an encoder error Homing: homing task is running Parked: not used PreCharge: completion of a home command when a preload is taken on the pump head Relaxed: motor H-Bridge is de-energized Running: pump in motion Stationary: idle Stopping: pump transitioning from running to stopped state UnHomed: not used
Relax Motors | Instructs the instrument to relax the specified motors; disabling the H-bridge FET drivers for each motor. Motor movements will be disabled. Ignores any parameters as if they weren't entered. | COMMAND OUTPUT | Motor Status Format: "%s" Example: "Relaxed"
Reset | Instructs the instrument to perform a soft reset after a specified delay (in ms). | COMMAND INPUT | Delay Time Units: ms Format: "%u" Range: 0 - 65535
Reset Settings | Reset the pump GUI settings and restore everything to the default values. This is a debug only command and not supported in the field. | COMMAND INPUT
Restore Factory Defaults | Delete the pump settings which restores everything to the default values. Deletes all the methods on the device. This is a debug only command and not supported in the field. | COMMAND INPUT
Get Pump Time | Requests the amount of time that the pump has been running. The pump time is reset to zero each time the flow rate is set to zero, and therefore is not saved over power cycles. | COMMAND OUTPUT | Pump Time Units: minutes Format: "%.2f" Example: "0.25"
Set Maintenance Counter | Command to set the number of pump strokes for the pumping system. | COMMAND INPUT | Count Format: "%u" Range: >= 0 Notes: if not entered the value is automatically set to zero OPTIONAL
Get Maintenance Counter | Returns the total number of complete pump strokes taken. | COMMAND OUTPUT | Count Format: "%u"
Get NVM | Instructs the instrument to report the integer value of the specified NV memory index. Parameter Index Values: Vary depending on product-specific memory map definitions. See “NVM Details” tab for memory map. Reports the requested NV memory integer parameter. NOTE: This command can be sent before the instrument has been Locked | COMMAND INPUT | NVM Index Format: "%u" Range: See "NVM Address / Range" column in "NVM Details" tab Example: "1021"
|  | COMMAND OUTPUT | NVM Index Format: "%u" Range: Contact Gilson R&D Example: "1021" | NVM Value Format: "%u" Example: "2500"
Set NVM | Instructs the instrument to set the specified NVM parameter at the specified NVM address. | COMMAND INPUT | NVM Index Format: "%u" Range: Contact Gilson R&D Example: "1021" | NVM Value Format: "%u" Notes:values are unchecked, errors are not generated if user does not adhere to defined ranges Example: "2500"
Get NVM String | Instructs the instrument to report the specified NV memory string parameter (serial number, name, maintenance dates). Parameter Index Values: Vary depending on product-specific memory map definitions. NOTE: This command can be sent before the instrument has been Locked Reports the requested NV memory integer parameter. | COMMAND INPUT | NVM Index Format: "%s" Legend: "Serial#": instrument serial number "Name": instrument name "Config" instrument hardware configuration "PHinstall": pump head installation date (any date format) "CVinstall": check valve installation date (any date format) "PSinstall": piston seal installation date (any date format)
|  | COMMAND OUTPUT | NVM Value Format: "%s" Range: "Serial#": 15 characters "Name": 39 characters "Config": 79 characters "PHinstall": 15 characters "CVinstall": 15 characters "PSinstall": 15 characters Notes: "Invalid parameter index" if error occurs
Set NVM String | Instructs the instrument to set the specified NV memory string parameter (serial number, name, maintenance dates). | COMMAND INPUT | NVM Index Format: "%s" Legend: "Serial#": instrument serial number "Name": instrument name "Config" instrument hardware configuration "PHinstall": pump head installation date (any date format) "CVinstall": check valve installation date (any date format) "PSinstall": piston seal installation date (any date format) | NVM Value Format: "%s" Range: "Serial#": 15 characters "Name": 39 characters "Config": 79 characters "PHinstall": 15 characters "CVinstall": 15 characters "PSinstall": 15 characters Notes: "Invalid parameter index" if error occurs
Reset NVM | Instructs the instrument to reset its NVM back to default | COMMAND INPUT | Mode Format" "%u$ Range: "0" - Resets all to defaults but: Instrument Name, Hardware Configuration (Controller Board), Pump Board Name, Hardware Configuration (Pump Board), Serial Number "1" - Resets all to defaults but: Instrument Name, Pump Board Name, Serial Number
Boot Sequence Complete | Sent after the ICI boots and gets the versions from the controller and pump module boards | UNSOLICITED EVENT
Dispense Complete | Sent after Dispense by Time or Dispense by Volume completes, is stopped (Stop Pump, Flow Rate set to 0 or an error occurs) | UNSOLICITED EVENT | Dispensed Volume or Time Units: mL or minutes Format: "%.2f" Example: "5.34"
Pressure Sample | Reports a series of time stamped pressure samples. Maximum of 10 samples per report. | UNSOLICITED EVENT | Timestamp¦Pressure Units: ms ¦ bar Format: "%u¦%.2f" Example: "64000¦500.25" Notes: each sample needs parsed into the time and pressure and is delimited by a pipe | Timestamp¦Pressure Units: ms ¦ bar Format: "%u¦%.2f" Example: "64000¦500.25" Notes: each sample needs parsed into the time and pressure and is delimited by a pipe OPTIONAL | Timestamp¦Pressure Units: ms ¦ bar Format: "%u¦%.2f" Example: "64000¦500.25" Notes: each sample needs parsed into the time and pressure and is delimited by a pipe OPTIONAL | Timestamp¦Pressure Units: ms ¦ bar Format: "%u¦%.2f" Example: "64000¦500.25" Notes: each sample needs parsed into the time and pressure and is delimited by a pipe OPTIONAL | Timestamp¦Pressure Units: ms ¦ bar Format: "%u¦%.2f" Example: "64000¦500.25" Notes: each sample needs parsed into the time and pressure and is delimited by a pipe OPTIONAL | Timestamp¦Pressure Units: ms ¦ bar Format: "%u¦%.2f" Example: "64000¦500.25" Notes: each sample needs parsed into the time and pressure and is delimited by a pipe OPTIONAL | Timestamp¦Pressure Units: ms ¦ bar Format: "%u¦%.2f" Example: "64000¦500.25" Notes: each sample needs parsed into the time and pressure and is delimited by a pipe OPTIONAL | Timestamp¦Pressure Units: ms ¦ bar Format: "%u¦%.2f" Example: "64000¦500.25" Notes: each sample needs parsed into the time and pressure and is delimited by a pipe OPTIONAL
Input Contacts Event | Reports the current states of the digital inputs when at least one input has changed. | UNSOLICITED EVENT | Start/Stop Format: "%s" Range: "Open", "Closed" | Pause/Resume Format: "%s" Range: "Open", "Closed" | Program Wait Format: "%s" Range: "Open", "Closed" | Error Format: "%s" Range: "Open", "Closed" | Power Down Format: "%s" Range: "Open", "Closed"
Error Event | Reports an error when one occurs that is not a result of command execution. | UNSOLICITED EVENT | Error Message Format: "%s" Example: "Encoder Error"
Pump Run Status | Provides updated values for the data displayed during a run. Sent unsolicited to the host while in remote mode, and to the GUI_APP while running. | UNSOLICITED EVENT | Current Flow Rate Units: mL/min Format: "%.3f" Example: "10.000" | Current Pressure Units: bar Format: "%.2f" Example: "500.25" | Current Solvent Level Units: mL Format: "%.3f" Example: "10000.25" | Current Run State Format: "%s" String Values: "Running" "Run Complete" "Paused" "Stopped" | Current Error Condition Format: "%s" String Values: "High Pressure" "High Pressure Method" "Over 75 Pressure" "Low Pressure" "Low Pressure Method" "Low Volume" "Input 4 Method" "Normal" | Method Name Format: "%s" Range: 16 characters | Current Method Loop Format: "%u" Range: 1 - 10000 | Current Method Event Number Format: "%u" Range: 1 - 50 | Pause Timer Time Units: minutes Format: "%.3f" Example: "5.314" | Current Timestamp Units: minutes Format: "%.3f" Example: "5.314" | Ramp Time Units: minutes Format: "%.3f" Example: "5.314" | Method Loop Time Units: minutes Format: "%.3f" Example: "5.314" | Current Dispense Time or Volume Units: minutes or mL Format: "%.3f" Example: "5.314" | Operation Hours Units: hours Format "%.2f" Example: "5.34"